## Slide 1: AFP-L3 and DCP Biomarkers for Liver Cancer
March 7, 2017
Hiro Yamada
Wako Diagnostics

## Slide 2: Outline
Part 1
Introduction of Wako Hepatocellular Carcinoma
(HCC) biomarkers
Part 2
Experiences with approaching the FDA for the HCC
biomarkers
Part 3
Current clinical study data
Notes: My name is Hiro Yamada, Senior Clinical Affairs Manager of Wako Life Sciences. Wako is an international company based in Osaka Japan, and its American head office in Mountain View, California.
Today I will present to you our HCC biomarkers, AFP, AFP-L3 and DCP on Wako’s immunoassay systems and some clinical study data on early detection of hepatocellular carcinoma (HCC).

## Slide 3: AFP-L3 (Lectin reactive alpha-fetoprotein)
LCA ； Lens culinaris agglutinin
AFP-L3% = AFP-L3 / (AFP-L1+AFP-L3) x 100
Carbohydrate chain
AFP-L1
Gal
Sia
Man
GlcNac
Fucose
AFP-L3
AFP-L3
Notes: This is one of our HCC biomarkers, Lectin reactive alpha-fetoprotein, which is called “AFP-L3”. AFP-L3 has the structure in which fucose is attached to the carbohydrate chain in AFP-L1. Based on the differences of their structures, AFP-L1 and AFP-L3 are detected separately and the percentage of AFP-L3 / AFP is calculated.

## Slide 4: DCP (des-gamma-carboxy prothrombin)
Glutamic acid or γ-carboxyglutamic acid
Amino acid
37- 579
NH2
1
10
30
20
Notes: This is another HCC biomarker, des-gamma-carboxy prothrobin, which is called “DCP”. DCP is abnormal prothrombin which has different structure including glutamic acid or gamma-carboxyglutamic acid shown as symbols of red triangle. (This is also called as “PIVKA-II”).

## Slide 5: Immunoassay Systems for HCC Biomarkers
Lectin Affinity Electrophoresis
Electrophoresis. 1989 Aug-Sep;10(8-9):562-7
2nd Generation: LiBASys
AFP-L3: May 2005
DCP : Jan. 2007
3rd Generation:
μTASWako i30
AFP-L3 / DCP:
Feb. 2011
In the US market

## Slide 6: Outline
Part 1
Introduction of Wako Hepatocellular Carcinoma
(HCC) biomarkers
Part 2
Experiences with approaching the FDA for the HCC
biomarkers
Part 3
Current clinical study data

## Slide 7: Commercialization Steps of LiBASys
Column
Analyzer
Cuvettes

## Slide 8: Commercialization Steps of LiBASys
C
C

## Slide 9: Submission of AFP-L3 on LiBASys
Wako LBA AFP-L3 Test Kit
Wako AFP-L3 Calibrator set and Control Set
Wako LiBASys Analyzer and accessories
Submission type: 510(k) De Novo
Substantial Equivalence: No predicate
2004 July

## Slide 10: Non-Clinical Data
Data Summary
Non-clinical Data
Sensitivity (LOD)
Within-run precision
Total precision
Accuracy
Linearity
Specificity (Interference studies)
Stability
Long term stability
Stability after reconstitution
Open vial stability

## Slide 11: Clinical Study Data
HCC | HCC DX | Non-HCC | Suspicious | Total
HCC →HCC | CH/LC →HCC | CH/LC →CH/LC | CH/LC →Suspious | 
54 | 39 | 273 | 71 | 437
Multi-center
Prospective
Surveillance Study
Lahey
VCU
Miami
Mt. Sinai
Toronto
UCSF
UPenn
Cut-Off Value
Sensitivity
Specificity
Positive period before diagnosis
Clinical course
Relative risk

## Slide 12: Submission of AFP-L3 on LiBASys
2005 May
Cleared by FDA as Class II
Intended for use: AFP-L3 is intended for in vitro diagnostic use as aid in the risk assessment of patients with chronic liver disease for development of hepatocellular carcinoma in conjunction with other laboratory findings, imaging studies and clinical assessment.
Continued to communicate with FDA
Q&A with FDA

## Slide 13: Publication of AFP-L3 on LiBASys

## Slide 14: Submission of DCP on LiBASys
Wako LBA DCP Test Kit, Calibrator set and Control Set
Submission type: Traditional 510(k), K062368
Substantial Equivalence: Wako LBA AFP-L3
2006 August
2007 January
Cleared by FDA as Class II
Intended for use: DCP is intended for in vitro diagnostic use as aid in the risk assessment of patients with chronic liver disease for development of hepatocellular carcinoma in conjunction ……
Q&A with FDA

## Slide 15: μTASWako i30 system
50 mm
17 mm
Microfluidic chip
Analyzer

## Slide 16: Integrated Immunoassay with Microfluidic Chip
Dispensing Mixing Washing Separation Detection
High speed, high precision
and high sensitivity
9 min for first results ,
2 min per subsequent results
25 test per hour
Notes: The i30 platform integrates all steps of immunoassay including dispensing, mixing, washing, separation, and detection, being carried out on chip. The i30 analytical microfluidics systems couples isotachophoresis and lectin affinity capillary electrophoresis on a single fully automated platform.

## Slide 17: Submission of μTASWako i30 system
Submitted devices:
Submission type: Traditional 510(k)
510(k) number: K100464
Predicate device:
μTASWako AFP-L3, Calibrator Set, Control L and H
μTASWako DCP, Calibrator Set, Control L and H
μTASWako i30 system
AFP-L3 test system including calibrators and controls was previously cleared as LBA AFP-L3 on LiBASys
DCP test system including calibrators and controls was previously cleared as LBA DCP on LiBASys
2010 February

## Slide 18: Submission of μTASWako i30 system
2011 February
Cleared by FDA as Class II
Intended for use: In DCP is intended for in vitro diagnostic use as aid in the risk assessment of patients with chronic liver disease for development of hepatocellular carcinoma in conjunction ……
Q&A with FDA

## Slide 19: Outline
Part 1
Introduction of Wako Hepatocellular Carcinoma
(HCC) biomarkers
Part 2
Experiences with approaching the FDA for the HCC
biomarkers
Part 3
Current clinical study data

## Slide 20: Inclusion of AFP-L3 & DCP Improves Detection of Early HCC

## Slide 21: Inclusion of AFP-L3 & DCP Improves Detection of Early HCC

## Slide 22: Case 1 – Risk assessment of HCC development by AFP-L3
AFP-L3
Normal: <10 %
AFP
Normal: <20 ng/mL
DCP
Normal: <7.5 ng/mL
Level of Biomarker
Cut-off
Level
13.8%
34.6%
87.7%
69.7%
Only AFP-L3 increased before HCC diagnosis by MRI.
------------------------------------- Elevation of AFP-L3 indicates higher risk of developing HCC and can be a trigger to do CT/MRI for early detection.
Hann HW, et al. J Med Microb Diagn 2014; 3: 130.
Months Before Diagnosis by imaging
MRI dx of HCC

## Slide 23: Case 2 – Risk Assessment Using AFP-L3 and DCP Elevation
Hann HW, et al. J Med Microb Diagn 2014; 3: 130.
Months Before Diagnosis by imaging
MRI dx of HCC
AFP-L3 and DCP were increased before HCC detection in some patients who had a low (normal) AFP level.
-----------------------------------------------------Adding AFP-L3 and DCP as risk markers to conventional surveillance tools such as AFP and ultrasound can increase chances of early detection of HCC.
0.30 ng/mL
0.32 ng/mL
17.65 ng/mL
7.4%
7.8%
38.6%
AFP-L3
Normal: <10 %
AFP
Normal: <20 ng/mL
DCP
Normal: <7.5 ng/mL
Cut-off
Level
Level of Biomarker

## Slide 24: Combined Biomarker Testing
AFP-L3 Positive
15 (20.3%)
DCP Positive
3 (4.1%)
AFP Positive
8 (10.8%)
7
(9.5%)
15
(20.3%)
4
(5.4%)
15
(20.3%)
All Negative
7
(9.5%)
Data collected for FDA submission
If HCC biomarkers are used prior to imaging:
In this study of 74 patients, the use of AFP alone would have detected 45 HCC cases and missed 29 cases of HCC
 | Cut off | Positive
AFP | ≥ 20 ng/mL | 45 (60.8%)
AFP-L3 | ≥ 10% | 49 (66.2%)
DCP | ≥ 7.5 ng/mL | 29 (39.2%)
All combined | Same as above | 67 (905%)

## Slide 25: HCC Biomarker Models
Z = -10.08 + 1.67 x [Gender] + 0.09 x [Age] + 0.04 x [AFP-L3]
+ 2.34 x log[AFP] + 1.33 x log[DCP]
= Gender, Age, AFP-L3, AFP, and DCP
GALAD Score for Diagnosis of HCC

## Slide 26: International Validation of GALAD

## Slide 27: International Validation of GALAD

## Slide 28: GALAD Evaluation using EDRN Phase 2 Cohort

## Slide 29: Canadian RCT
This study expects to demonstrate that addition of biomarkers
will increase the detection rate by at least 10%.

## Slide 30: VA Houston Prospective Study

## Slide 31: Thank you for your attention!