## Slide 1: Biomarker Verification and Clinical Grade Assays
Sanford A. Stass, MD
University of Maryland Baltimore
Biomarker Reference Laboratory
32nd EDRN Steering Committee Meeting September 12-14, 2017

## Slide 2: Clinical Laboratory Improvement Amendments (CLIA)
CLIA-88(final regulations published in 1992): Concerns arose regarding quality of laboratory testing (cholesterol screening, Pap smears).
Law requires laboratories doing testing for “diagnosis, prevention, or treatment of any disease or impairment of, or the assessment of the health of human beings” be certified (initially by HCFA, now by Center for Medicaid and Medicare Services (CMS).

## Slide 3: Clinical Laboratory Improvement Amendments (CLIA) (Cont’d)
The Clinical Laboratory Improvement Amendments (CLIA) (est. 1988, finalized 1992, and revised 2003).
  CMS regulates all laboratory testing through CLIA
  CLIA addresses issues of good laboratory practice, such as:
    Equipment maintenance
    Proficiency testing
    Competency testing
    Personnel training and documentation
    Result reporting
  CLIA does NOT address clinical utility, test sensitivity, result interpretation

## Slide 4: College of American Pathologists (CAP)
World’s largest pathologists association.
Leader in laboratory quality assurance.
Federal government recognizes CAP as equal to or more stringent than the government’s own inspection program.
Deemed status with CLIA.
CAP accreditation awarded (7,585 US and 381 International laboratories) by Commission Laboratory Accreditation based on results of stringent on site inspection every two years including examination of:
  Standard operating procedures/laboratory records
  Quality control procedures/quality assurance programs
  Qualification of directors and staff
  Laboratory equipment
  Facilities
  Safety Program
  Overall laboratory management

## Slide 5: Pre-analytic errors
Collecting sample from wrong patient
Collecting the wrong sample
Mislabeling or failing to label sample
Storing sample incorrectly prior to testing
Improper Transportation of samples
Damaged reagents or test kits
What Can Cause Inaccurate Results?
2017 College of American Pathologists
CAP and Accreditation Overview 2017

## Slide 6: Analytic errors
Failing to follow an established algorithm
Reporting results when QC tests out of range
Incorrect sample or reagent measuring
Using improperly stored or expired reagents
Instrument calibration bias
Reporting results beyond acceptable linearity range
Conducting testing in wrong test mode
What can Cause Inaccurate Results?
2017 College of American Pathologists
CAP and Accreditation Overview 2017

## Slide 7: Laboratory Errors Happen
Root Causes?
  Training
  Competence
  Environment
  Policies and Procedures
  Test system performance
  Understandable report created with an effective interpretation
Data suggests that 24-30% of laboratory errors have an effect on patient care while actual or potential patient harm occurs in
3-12%. 1
Hawkins R. Managing the Pre- and Post-analytical Phases
of the Total Testing Process. Ann Lab Med 1012;32:5-16.
2017 College of American Pathologists
CAP and Accreditation Overview 2017

## Slide 8: CAP’S Accreditation Program helps laboratories achieve these outcomes using:
21 discipline specific checklists that are a blueprint for running a high quality laboratory.
  ~ 40% of requirements exceed federal requirements and reflect the more stringent standards of CAP.
Peer inspection model focuses on education and information sharing.
  Specialty inspectors for emerging high complexity areas
Proficiency testing monitoring beyond regulated analytes because all tests should be done correctly
2017 College of American Pathologists
CAP and Accreditation Overview 2017

## Slide 9: Standards and Checklists
Purposes:
  Standards are the broad principles the laboratory must meet in order to achieve accreditation.
  Checklists provide detailed requirements that inspectors use to determine whether laboratories meet the Standards
2017 College of American Pathologists
CAP and Accreditation Overview 2017

## Slide 10: CAP 2-year accreditation cycle covers the entire lab
Application
Onsite Inspection
Report Inspection Results to CAP
Deficiency Responses
CAP Review of Deficiency Responses
Accreditation Decision
Self Inspection
2017 College of American Pathologists
CAP and Accreditation Overview 2017
Notes: Accreditation is a continuous process by which the CAP assesses the laboratory.
CAP incorporates compliance assessment and process improvement into each step of the two-year accreditation cycle.
With submission of the Request for Application form, CAP will send you:
    The Application in which you will provide an overview of your laboratory
    Complete set of checklists to help you formally describe the services your institution and your laboratory provides
STEP #1: The laboratory completes and sends application to CAP
  Information on the application drives accreditation fees, the PT you will be required to conduct for continuous monitoring, your custom checklist, and the number of people and skill set needed for you inspection team.
Application also includes information on the services your institution provides (e.g., Cardiac Center, Organ Transplant) to patients, and the corresponding test menu and the structure and organization of your lab.
From receipt of the application, the laboratory has a maximum of one year to prepare. Along the way, they can ask CAP questions, consult with other accredited laboratories, and participate in CAP educational webinars and seminars occasionally provided
During this time, the laboratory must also demonstrate at least 6 months of PT performance on all analytes included in their activity menu (usually 2 shipments)
The Entire Lab
  CAP believes that laboratory accreditation provides the patient assurance that ALL processes and procedures are meeting requirements; consumers don’t know whether only one section of a laboratory is accredited—therefore, CAP does not segment its accreditation.
Activity Menu
  The laboratory provides information about its scope of testing and lists all reportable assays in the Activity Menu. The information provided is critical, as it will be used to customize checklists, to determine disciplines for which accreditation is granted.
Customized checklists
  CAP determines checklist usage from the activity menu completed for each laboratory section.
  Depending on the organization of the lab, more than one checklist may apply to any one laboratory section. In the example on the slide, the Chemistry section unit also performs urinalysis testing and therefore has both checklists.
  Checklists are customized based on the laboratory’s activity menu.
  Subdiscipline sections and other significant groups of requirements not pertinent to the testing performed in the laboratory will not be included.
  Customized checklists greatly reduce the number of non-applicable checklist requirements, thereby streamlining the inspection process.

## Slide 11: Standard Operating Procedures (SOPs) CLIA/CAP Compliance
Specimen Handling/Specimen Rejection Criteria/Specimen Storage
Principle/Procedure
Equipment
Supplies
Controls/Reagents/Reagent Preparation
Quality Control/Criteria for Controls
Interpretation and Procedure Limitations
Troubleshooting Guide
References

## Slide 12: Specimen Handling
Specimen collection and handling:
  Methods for collection
  Proper labeling
  Method for delivery
  Method for specimen preservation
  Procedure for safe handling
Proper recording of specimens received

## Slide 13: Specimen Handling (cont’d)
Criteria for rejection.
No aliquot is ever returned to the original container.
Aliquoting procedure prevents any possible cross-contamination.
Schedule for retaining specimens.
Verify identity and integrity through all steps.
Condition documented upon receipt (evidence of tampering, inadequate volume, etc.)
Specimens limited access, secured area.

## Slide 14: Equipment Validation/Quality Control (QC)
New equipment “clinically validated”.
All equipment regular Q.C.
  Q.C. records reviewed monthly.
  Records maintained in Q.C./Troubleshooting Notebooks.
SOP for set-up/normal operation of equipment.
Instructions/regular schedule for checking critical operating characteristics.
Function checks to detect trends/malfunctions.
Tolerance limits for acceptable function.
Instructions for troubleshooting/repairs.
Surveys to include UV/VIS spectrophotometer and analytical balance.

## Slide 15: Equipment Validation/QC (cont’d)
Document repairs/service procedures/maintenance.
Spectrophotometer wavelength calibration checked regularly.
All calibration curves rerun/verified after servicing or equipment recalibration.
Criteria for acceptable background levels taken regularly.

## Slide 16: Equipment Validation/QC (Cont’d)
Temperatures checked/recorded.
  Water baths,
  Heating blocks,
  Incubators and ovens,
  Refrigerators and freezers
Individual wells (or a representative sample thereof) of thermocyclers checked for temperature accuracy pre and post placed into operation.

## Slide 17: Miscellaneous Supplies
Volumetric glassware of certified accuracy (Class A, NIST Standard or equivalent).
  Non-certified volumetric glassware must be checked for accuracy of calibration before initial use and at specified intervals.
  Reconstitution of lyophilized calibrators, controls or proficiency testing materials, or any other tasks requiring accurate volumetric measurement, performed with measuring devices of Class A accuracy, or those for which accuracy has been defined and deemed acceptable for the intended use.
Appropriate thermometric standard device (e.g., NIST certified) for calibration checks.
All non-certified thermometers checked against an appropriate thermometric standard device.

## Slide 18: Controls/Reagents/Reagent Preparation
Document QC of reagent shipments of the same lot number
New reagent lots checked against old reagent lots or with suitable reference material before or concurrently when placed in service.
Appropriate validation studies performed/documented for use of any reagents that do not follow manufacturer’s recommendations.
Reagents/solutions properly labeled:
  Content and quantity, concentration or titer
  Storage requirements
  Date prepared or reconstituted by laboratory
  Used within labeled expiration date.

## Slide 19: Quality Control/Quality
All SOPs set specific criteria for controls.
  Tolerance/acceptability limits defined for all control procedures, control materials, and standards.
All assays include external positive, negative, and amplification controls as QC for the specific assay.
  Distinguish a true negative result from false negative due to PCR failure. False negatives are often due to PCR inhibitors in the specimen; demonstrate that a different DNA/RNA sequence from the one being tested is amplifiable in the same specimen.
Evidence of corrective action when control results exceed defined tolerance limits.
  Unacceptable controls documented on worksheet/Troubleshooting Notebook.
Quantity/quality (intactness) high molecular weight DNA/RNA assessed.
  Nucleic acids processed promptly/stored adequately to minimize degradation.

## Slide 20: Interpretation of Results
Criteria for assay parameters/data interpretation.
Clearly define interpretation guidelines resulting in a level of consistency between laboratories and clear delineation of positive, negative, or indeterminant results.

## Slide 21: Personnel Training & Competency
Training new staff.
Annual training/evaluation includes:
  Annual competency tests using in-house blind samples, CAP/CDC survey/exchange samples, or written tests.
  A minimum of 8 continuing education hours.
  HIPAA, Compliance, and Safety/Hazard training.

## Slide 22: Experience with BDL to UMB-BRL Assay Development and Implementation – Lessons Learned

## Slide 23
This assay consisted of Short Tandem Repeat PCR analysis of 15 previously validated microsatellite markers located within 14 gene loci (D4S243, D21S1245, FGA, D17S695, D16S476, D9S171, IFN-A, D20S48, D17S654, D16S310, TH01, D9S162, D9S747, MBP and MBPa).
Implementation issues between commercial and UMB-BRL CLIA/CAP laboratories.
  Difference in Primer pairs used
  Need for new Positive Control
  Refinement of SOP with regard to interpretation and acceptance criteria
  New reagents (PCR master mix and HLPC primers)
  Clear definition of interpretation guidelines that result in level of consistency between laboratories
  Clear definition of all aspects of these assay parameters in the SOP
MSA Assay Validation for
Detection of Bladder Cancer
SOP/Qualification

## Slide 24: Assay Development/Validation of GP73 Biomarker for Hepatocellular Carcinoma
Glycoproteomic studies suggest hepatocellular carcinoma has greatly elevated serum levels of “total” and fucosylated forms of normal Golgi resident glycoprotein (GP73) compared to cirrhosis.
Collaboration with Dr. Timothy Block (Drexel University) established ELISA for quantitative GP73.
Confirmed reagents and materials for the quantitative GP73 assay and established SOP with CLIA/CAP standards for assay in the UMB-BRL using LiCor equipment.
Parallel testing for the quantitative GP73 assay with Dr. Block’s research laboratory established that the ELISA assay demonstrated inter-laboratory (100% concordance) using the LiCor equipment, indicating successful transfer of the assay from Dr. Block’s laboratory to the UMB-BRL CLIA/CAP laboratory.
The Licor would not be the most likely equipment used in the clinical laboratory. LiCor optics is not as sensitive resulting in false negatives. The Liaison, or a like system that uses chemoluminescent technology, is a better fit for clinical laboratory since it is more sensitive and would result in fewer false negatives.

## Slide 25: Application of a Vimentin Gene Methylation Potential Biomarker for Colon Cancer
ABI 7900 HT
Laser dye excitation
Less sensitive/stable optics
BioRad CFX96
LED dye excitation
Sensitive, solid state digital optics
Solution
  UMB-BRL purchased BioRad CFX96 equipment
Transfer of Vimentin Gene Methylation Assay from BDL (Sanford Markowitz, MD, PhD, Case Western Reserve University) to UMB-BRL CLIA/CAP Laboratory
Methylation specific quantitative real-time PCR (qRT-PCR)
Assay Pre-validation Study Issues
Investigation of difference in results between BDL and UMB-BRL: UMB-BRL sent technologist to BDL to process the same set of samples side-by-side using the same SOP, reagents, and equipment.
The BDL and UMB-BRL staff obtained identical results for this exercise.
Conclusion: These differences likely due to the sensitivity of the equipment.

## Slide 26: Potential Biomarker for Colon Cancer Transfer of Vimentin Gene Methylation Assay From BDL (Case Western Reserve University) to UMB-BRL CLIA/CAP Laboratory (Cont’d)
PCR Amplification Buffer affects sensitivity
  Roche buffer (Essential) vs. ABI Taqman universal buffer.
    The Roche buffer produced lower Ct values (mid –high 20’s to low 30’s) which improves assay sensitivity vs Taqman (high 30’s to 40’s). Therefore, the Roche buffer was used.
Reagents (e.g. buffers) were verified, optimized and standardized for clinical application.
The UMB-BRL implemented HPLC purified primers and probes.
Certain elements of the assay were modified to adapt to CLIA/CAP standards
  Inclusion of standard curve using CpGemone DNA, etc.
  The UMB-BRL implemented use of diluted CpGenome DNA (methylated and unmethylated) versus diluted cell line.
Evaluation of DNA clean up on the stool samples
Repeat testing of previous pre-validation study samples demonstrated a high degree of concordance between BDL and UMB-BRL.
Specimen processing SOP for samples received at UMB-BRL developed.

## Slide 27: Overview Implementing BDL Laboratory Assays into CLIA/CAP-Compliant Assays Issues/Lessons Learned
“Paper Towel Syndrome”
  Lack of organized SOP without clear definition of all assay parameters (e.g. fragmented SOP written on paper towels, isolated notes in lab coat pockets, etc.)
Modification of SOPs
  BDLs may have used assays prior to discovery/development of a biomarker. SOPs may have been modified on an ongoing basis without documentation/rationale making it difficult to implement/evaluate assay efficacy.
Consistency of assay performance
  Alternate methods/different technologists
Staff training
    Appropriate training
    Potential for inconsistencies in running assay from batch to batch
    Assure staff complies with procedures as delineated in SOP and assure different staff do not “do it their own way”

## Slide 28: Overview Implementing BDL Laboratory Assays into CLIA/CAP-Compliant Assays Issues/Lessons Learned (Cont’d)
Compliance with CLIA/CAP QC Standards
  Standard controls or presence of controls in assay.
  Equipment calibration, SOP, records, validation, QC, maintenance and monitoring, including pipet QC
  Reagent QC including, source, tracking by lot and shipping number, storage conditions; HPLC grade not used
  Kits used for testing not at clinical level (e.g. kits purchased outside U.S. which may not meet clinical standards) and contents mixed
  QC of in-house reagents and expiration dates
  The same model of equipment/platform or different equipment may not function the same and produce comparable results between the developmental laboratory and the BRL.
Alternate methods used in the BDL for the same assay by different technologists or same assay performed differently by technologists in BDL makes it difficult for BRL to develop CLIA/CAP SOP.
Lack of clarity on how BDL calculates results
Reagents not stored at proper temperatures; temperatures of refrigerators/freezers not recorded
BDL assays may use equipment that is not in general use in a clinical laboratory
Budget comparisons between BDL and BRL may be different due to differences in final BRL assay.
Interpretation guidelines/results must be developed based on validation and be clearly defined

## Slide 29: BRL and BDL CLIA/CAP Assay Recommendations
If you don’t know where you are going you might end up some place else
Yogi Berra

## Slide 30: BRL and BDL CLIA/CAP Assay Recommendations
BDL and BRL communicate early in the process of assay development/implementation for testing in a CLIA/CAP laboratory.
Perform unblinded parallel studies between BDL and BRL during assay development/implementation to assess efficacy of all assay parameters.

## Slide 31: BRL and BDL CLIA/CAP Assay Recommendations (Cont’d)
Clearly validate and define interpretation guidelines that result in consistent agreement between BDL and BRL
Early identification by the BRL of promising platform(s) used by BDLs so BRLs will be prepared for technology transfer
Assays used by BDLs may not be in use in BRLs/clinical laboratories or are assay adaptations. BRLs need to learn and optimize reliability of these assays early in the process
Clearly define all aspects of assay parameters in the SOP
Continued parallel testing as the biomarker development laboratory develops and the BRL evaluates biomarker(s) for clinical implementation
Incorporate well defined quality assurance oversight/methods to provide more consistent results throughout the development and implementation process

## Slide 32
UMB-BRL Core Facilities and Resources
Potential Areas of Collaboration with Other Members of EDRN
Extensive laboratory technology and testing experience including translation of biomarkers
research into direct CLIA/CAP clinical laboratory testing in a stringent environment required
by regulatory agencies (federal and state) to assure quality patient testing.
Laboratories of Pathology CLIA/CAP/ASHI/FDA Accredited | Anatomic and Clinical Pathology Sanford Stass, MD | University of Maryland Medical Center
Pathology/Medicine Laboratory CLIA/CAP Accredited | Translational Genomics Laboratory Linda B. Jeng, MD, PhD | UMB, UMSOM
Genomic Technologies Core Facilities and Resources | Institute for Genome Sciences Claire M. Fraser, PhD | UMSOM
 | Biopolymer/Genomics Core Facility Nicholas P. Ambulos, Jr., PhD | UMSOM
Information Technology, Informatics and Biostatistics | Biostatistics Core Service Søren M. Bentzen, PhD, DMSc | UMSOM
Pathology Laboratory CLIA/CAP Accredited | Molecular Diagnostics Laboratory Richard Y. Zhao, PhD | UMSOM
Pathology Laboratory CLIA/CAP Accredited | Clinical Pathology and Research Laboratory Robert Christenson, PhD | UMSOM
Pathology Laboratory CLIA/CAP Accredited | Cytogenetics Laboratory Ying Zou, MD, PhD | UMSOM
Pathology Laboratory CLIA/CAP and ASHI Accredited | Flow Cytometry Core Laboratory Immunogenetics Laboratory Debra KuKuruga, PhD | UMSOM
University of Maryland Baltimore (UMB), University of Maryland School of Medicine (UMSOM) and Off-Campus Core Facilities and Resources

## Slide 33
UMB-BRL Core Facilities and Resources
Potential Areas of Collaboration with Other Members of EDRN
University of Maryland Baltimore (UMB) and Off-Campus Core Facilities and Resources
Structural Biology/Mass Spectrometry/Proteomics | BIACore Core Facility Robert Bloch, PhD | UMSOM
 | Center for Mass Spectrometry Resources Maureen A. Kane, PhD | University of Maryland School of Pharmacy (UMSOP)
 | Nuclear Magnetic Resonance Center David J. Weber, PhD (Dir.) Kristen M. Varney, PhD (Manager) | UMSOM and UMSOP
 | X-ray Crystallography Core Facility Eric Sundberg, PhD | UMSOM
Drug Development Core Facilities and Resources | Computer Aided Drug Design Center Alex MacKerell, PhD | UMSOP
 | Center for Nanomedicine and Cellular Delivery Peter Swann, PhD | UMSOP
Imaging Core Facilities and Resources | Center for Fluorescence Spectroscopy Joseph Lakowitz, PhD | UMSOM
 | Confocal Microscopy Core Facility Thomas Blanpied, PhD | UMSOM
 | Electron Microscopy Core Facility Ru-ching Hsia, PhD | UMSOM and UMSOP
 | Center for Translational Imaging Rao Gullapalli, PhD | UMSOM
 | Magnetic Resonance Research Center Rao Gullapalli, PhD | UMSOM
Pathology Laboratory | Tissue Bank Core Laboratory Olga B. Ioffe, MD | University of Maryland Marlene and Stewart Greenebaum Comprehensive Cancer Center

## Slide 34
UMB-BRL Core Facilities and Resources
Potential Areas of Collaboration with Other Members of EDRN
University of Maryland Baltimore (UMB) and Off-Campus Core Facilities and Resources
Mid-Atlantic Shared Services Consortium | Metabolomics & Glycomics Amrita Cheema, PhD Radoslav Goldman, PhD | Georgetown Lombardi Comprehensive Cancer Center
 | Analytic Pharmacology Michelle Rudek, Pharm D, PHD Cell Processing and Gene Therapy Facility M. Victor Lemas, PhD | Sidney Kimmel Comprehensive Cancer Center at Johns Hopkins
 | Biomolecular Magnetic Resonance Facility Jeff Ellena Biomedical Mass Spectrometry Laboratory: Proteomics Nicholas Sherman Molecular Imaging Core Stuart Berr Flow Cytometry Facility Joanne Lannigan Lymphocyte Culture Center: Monoclonal Antibody Production Bill Sutherland | University of Virginia Cancer Center